## Slide 1: 真空腔體系統的設計、控制與操作

## Slide 2: 真空系統的建置，須考量如下
主要思考項目
真空壓力範圍
製程方式
抽氣方式
成本
操作及日後維護方便
次要思考項目
適切的真空法蘭面
合適的真空計的選用
適切的真空幫浦的選用
適切閥門選用
材料選用
腔體大小形狀
溫度
氣 電 引入
安全
容易安裝 容易維修 容易操作
操控系統
紀錄
以上主要＆次要項目 相戶影響 、牽制 才有結論

## Slide 3: 真空系統含那些主要項目及操作方式
真空項目
Vacuum Chamber
Vacuum Pump
Vacuum Gauges
Vacuum Valves
真空引入(氣,電…),
Sensor( Diode sensor,….)
機械手臂
Heater
Vacuum 元件( 視窗, Bellow,adapter…)
操作方式
完全目視/完全手動
完全目視/完全手動( 感應防呆)
完全目視/部分手動,部分自動 (感應防呆)
高程度自動化行程

## Slide 4: Chamber
大小
形體
強度
使用的真空壓力
溫度
腐蝕性?
引入
機械傳輸 轉動
未來 PM
圓柱/方形
壓力產生的形變( 1kg/cm2) & 腔們強度 , 載重( Valve, Pump......)
304, 316, 304ESR, 鋁合金......
O-RING/All Metal
冷卻水 gas 電
robot

## Slide 5
Flanges 多寡, Size, Oring or All metal ( CFF, ISO, KF, ANSI?)
Viewport?
Turbo pump/Cryo pump 採 橫掛或垂直
震動考量
feedthrough
Chamber

## Slide 6: 本公司及母公司的相關範例

## Slide 7: 本公司及母公司的相關範例

## Slide 8: Vacuum Levels
Pressure
10-12
100
1
10-4
10-6
10-8
10-10
10-2
200 (322)
100 (161)
80 (129)
60 (97)
40 (64)
20 (32)
Torr
1000 (1609)
400 (644)
miles (km)
760
Pa
104
100
10-2
10-4
10-8
1
105
10-10
10-6
0 Sea Level
Ultra-High
Vacuum
High
Vacuum
Rough
Vacuum
Typical
Base
Pressures
{
{
Process
Pressures
Altitude
}
}
Cirrus Clouds
Aircraft
}
Low-Earth
Satellites
}
Space Shuttle
Notes: Notes:
__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

## Slide 9: Vacuum Pumps
Pumping System Ranges
10-10 10-9 10-8 10-7 10-6 10-5 10-4 10-3 10-2 760
Cryopump
Oil Diffusion Pump with Cold Trap
Turbomolecular backed with 2 Stage Mechanical Pump
Mechanical
Pressure (Torr)
Ultra-High Vacuum
High Vacuum
Rough Vacuum
Waterpump
Ion Pump/TSP PUMP
Dry Scroll pump
Notes: Notes:
__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

## Slide 10: 如何選擇幫浦
(超)高真空幫浦
Turbo Pump
Diffusion Pump
Ion Pump/TSP/NEG
Cryo Pump
複合幫浦
Primary Pump
Pumping Speed ( 在不同壓力下的 抽氣速度)
Oil or Dry
腐蝕程度 粉塵程度
N2 Purge( 同上 或 氣爆)
節電

## Slide 11: CRYOPUMPING BASICS: Operation - Cryoadsorption
Air gases and water vapor still condensed - noncondensible gases captured.
80 K
<20 K
H2O
N2
Ar
Ne
H2
He
1st Stage
2nd Stage
Notes: Notes:
__________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

## Slide 12: GAS LOAD and Pump througtput
Another name for Q, the throughput, is gas load. Where does gas load, “Q” come from?
Outgassing
Leaks
Virtual
Real
Backstreaming
Diffusion
Permeation
(Don’t forget Volume gas!)
Pump Q(throughput)
Q pump>Q gas load total
Pressure goes down

## Slide 13: 選擇Valves
Gate Valve
Butterfly Valve
Roughing Valves( inline/right angle)
Needle valve
Ball valve
進氣閥
微漏閥
others
不同的閥門有其功能性 , 其中需考量：真空度、氣導、逸氣、開關次數、洩漏率、價錢 ......等。

## Slide 14: 選擇真空計

## Slide 15: Types of Vacuum Gauges
Direct Gauges (Displacement of a wall)
Solid Wall
Liquid Wall
Diaphragm
Bourdon Type
Capacitance Diaphragm
U-Tube Manometer
McLeod

## Slide 16: Types of Vacuum Gauges (cont.)
Charge Generation
(Ionization)
Energy Transfer
(Heat Loss)
Spinning Rotor
Indirect Gauges
(Measurement of a gas property)
Momentum Transfer
(Viscosity)
Thermocouple
Pirani
CONVECTRON
Cold Cathode
Bayard-Alpert
Triode
Schulz-Phelps
STABIL-ION
Hot Cathode
Inverted Magnetron
Penning

## Slide 17
Ranges of Vacuum Gauges
Pressure (Torr)
10-12
10-11
10-10
10-9
10-8
10-7
10-6
10-5
10-4
10-3
10-2
10-1
1
10
100
1000
Nude Bayard-Alpert
STABIL-ION
Tubulated Bayard-Alpert
Triode
Schulz-Phelps
Cold Cathode
Spinning Rotor
Thermocouple
Pirani
CONVECTRON
Capacitance Diaphragm
McLeod
U-Tube Manometer
Diaphragm
Bourdon

## Slide 18: 決定真空行程

## Slide 19: 範例一簡單版 :Leak Detectors Valves Logistics

## Slide 20: 範例二複雜版 (母公司範例)

## Slide 21: CTI-CRYOGENICS Compressor won’t run
Notes: Red = return pressure helium
Green = supply pressure helium
Blue = cooling water Note the the system is sensitive to water flow direction, as well as rate and temperature.
Yellow = oil flow

## Slide 22: 紀錄
日期 | 時間 | 毫秒 | Eq ID | S/N | Op ID | Test Result | Leak Rate | Leak Trend
2018/6/13 | 9:22:22 | 240 | 0 | 0 | 0 | 2 | 0.7E-09 | -8.93
2018/6/13 | 9:22:23 | 540 | 0 | 0 | 0 | 2 | 0.7E-09 | -8.929
2018/6/13 | 9:22:24 | 210 | 0 | 0 | 0 | 2 | 0.7E-09 | -8.928
2018/6/13 | 9:22:25 | 310 | 0 | 0 | 0 | 2 | 0.7E-09 | -8.926
2018/6/13 | 9:22:26 | 280 | 0 | 0 | 0 | 2 | 0.8E-09 | -8.92
2018/6/13 | 9:22:27 | 270 | 0 | 0 | 0 | 2 | 0.1E-09 | -8.99
2018/6/13 | 9:22:28 | 230 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:29 | 220 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:30 | 320 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:31 | 300 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:32 | 270 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:33 | 250 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:34 | 220 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:35 | 320 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:36 | 300 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:37 | 290 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:38 | 270 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999
2018/6/13 | 9:22:39 | 240 | 0 | 0 | 0 | 2 | 0.0E-09 | -8.999

## Slide 23: Endura® SL System Layout

## Slide 24: 控制系統需要控制讀取的元件
Turbo Pump
Ion Pump/TSP
Primary Pump
Hi /Low Vacuum Gauges
Sensors ( Position, Temperature….)
……..

## Slide 25: 軟體
Window,UNIX,DOS, LABVIEW…
AHA 採 PLC +HMI
AMAT/ Varian Implanter/ R&D
小型系統/遠端監控/Cost Saving

## Slide 26: AHA 五合一真空系統控制器
Ion pump
Turbo Pump
Roughing Pump
Vacuum Gauges
Valve Control
Other Control( Heater on/off, Power on/off)